## Sheet1
1 | ROB30001989 | 10309710 |  | 2025-05-05 | other |  |  |  | Not Available. | 2025-00117778-1：High Test Results | 2025-05-05 | medical organization | KAISER FOUNDATION HOSPITAL, UNITED STATES | The customer reported that falsely elevated High-Sensitivity Troponin I (TnIH) results were obtained on two patient samples on an ADVIA Centaur CP Analyzer. The initial results were reported to the physician(s). For the first patient, two new samples were drawn at different times and processed on the original analyzer, which recovered lower than the initial result. Both the second and third drawn sample results were considered correct. Then, the customer repeated the original sample on the original analyzer and obtained a similar lower result, which was considered correct. The result of 21.0 pg/mL was reported as the correct result to the physician(s). For the second patient, the sample was repeated on an alternate ADVIA Centaur XP Analyzer. The repeat result recovered lower than the initial result and was considered correct. The repeat result was reported to the physician(s). There are no known reports of patient intervention or adverse health consequences due to the falsely elevated TnIH results. Assay: High-Sensitivity Troponin I (TnIH) Units: pg/mL Date of Testing: 05-May-2025 Sample 1: Initial Result: 331 (Erroneous, Reported) 1st redrawn Sample Result: 21 (Correct) 2nd redrawn Sample Result: 21 (Correct) Repeat Result from the Original Sample: 21 (Correct, Reported) Sample 2: Initial Result: 130 (Erroneous, Reported) Repeat Result (Alternate Instrument): <3 (Correct, Reported) | yes | A United States (US) customer contacted a Siemens Customer Care Center (CCC) and reported that falsely elevated High-Sensitivity Troponin I (TnIH) results were obtained on two patient samples on an ADVIA Centaur CP Analyzer. Quality Control (QC) was within the acceptable range on the day of the event. A Siemens Customer Service Engineer (CSE) was dispatched to the customer's site. During the visit, the CSE performed the total service call, found a luminometer with dried liquid inside. The CSE cleaned the luminometer, verified the dark count, replaced the aspirate wash pump tubing and aspirate line, verified the sample and reagent probes calibrations, decontaminated the rinse and wash 1 lines, and ran QC 10 times and patient samples, which recovered within acceptable limits. Siemens further evaluated the event and concluded that the wash block area of the system potentially contributed to the falsely elevated TnIH results as unbound lite reagent not being fully washed from the cuvette. The instrument is performing according to specifications. No further evaluation of this device is required.
2 | CM00999 | 11067000 |  | 2025-05-14 | other |  |  |  | Not Available. | 2025-00120323-1：Low Test Results | 2025-05-14 | medical organization | HUSLAB HYVINKÄÄN SAIRAALA, FINLAND | The customer reported a falsely depressed Sodium (Na) result was obtained on a patient sample on an Atellica CH 930 Analyzer. The erroneous result was reported to the physician(s). The sample was repeated on two alternate Atellica CH 930 Analyzers. The repeat results were higher than the erroneous result. The first repeat result was reported, as the correct result, to the physician(s). There are no known reports of patient intervention or adverse health consequences due to the falsely depressed Na result. Date of Testing: 14-May-2025 Assay: Sodium (Na) Units: mmol/L Sample Identifier: HM253VPFP Initial Result: 114.4 (Erroneous, Reported) Repeat Result on an alternate Atellica CH 930 Analyzer: 127 (Correct, Reported) Repeat Result on an alternate Atellica CH 930 Analyzer: 128 (Correct) | yes | An Outside of the United States (OUS) customer contacted a Siemens Customer Care Center to report a falsely depressed Sodium (Na) result was obtained on a patient sample on an Atellica CH 930 Analyzer. The customer inspected the sample and found no abnormalities. Quality Controls (QCs) recovered within ranges on the day of the event. A Siemens Customer Service Engineer (CSE) was dispatched to the customer’s site. During this visit, the CSE ran a full Atellica Test Protocol (ATP). The dilution mixer test recovered low and the Coefficient of Variation (CV) was not in range for the dilution and sample probe. The CSE adjusted the dilution mixer and performed an auto alignment. The CSE replaced the dilution arm level sense cable, pressure sensor, pressure valves and sample probe. The dilution probe, sample probe and reagent arms were adjusted. Dilution probe and sample probe tests were performed and recovered within ranges. On a subsequent visit the CSE replaced the reagent mixer. The water bath was emptied and filled. The reagent mixer and reagent arms were adjusted, and mixer tests and QC were run, which recovered within ranges. The cause of the falsely depressed Na result is unknown. The instrument is performing according to specifications. No further evaluation of this device is required.
3 | DV331154 | 10284473 |  | 2025-05-16 | other |  |  |  | Not available. | 2025-00121197-1：Low Test Results | 2025-05-16 | medical organization | JOHN MUIR MEDICAL CENTER, UNITED STATES | Three samples gave falsely depressed Carbon Dioxide (CO2) test results on a Dimension Vista 500. The initial test results were released to the physician(s). The same samples were repeated on the same Dimension Vista 500 and corrected reports were released to the physician(s). There are no known reports of patient intervention or adverse health consequences due to the behavior. Event date: 16-May-2025. Sample ID: 4003672886, 72 year old male Test: Carbon Dioxide, mmol/L Initial result: 14 (discordant, reported) Repeat result: 23 (correct, reported) Sample ID: 4003672825, 60 year old male Test: Carbon Dioxide, mmol/L Initial result: 16 (discordant, reported) Repeat result: 25 (correct, reported) Sample ID: 4003672827, 72 year old male Test: Carbon Dioxide, mmol/L Initial result: 10 (discordant, reported) Repeat result: 25 (correct, reported) | yes | Three samples gave falsely depressed Carbon Dioxide (CO2) test results on a Dimension Vista 500. A Siemens Operational Support (SOS) specialist connected remotely to the customer site and determined all the impacted tests were on Server 1. A Siemens Customer Service Engineer (CSE) was dispatched to the customer site to inspect the Dimension Vista 500. The CSE replaced the R1 mixer and replaced and aligned the R1 and R2 probes. The CSE also replaced the sample drain and primed the sample probe. Quality control materials were run with acceptable results. The cause of the discordant CO2 and BUN test results is unknown. The analyzer is performing within specifications. No further evaluation of this analyzer is needed.
4 | 62315109 | 10997841 | 2026-01-15 | 2025-05-20 | other | years | 71 | Male | No additional patient information was provided. | 2025-00121334-1：High Test Results | 2025-05-20 | medical organization | Kliniken Landkreis Heidenheim, GERMANY | The customer observed an elevated Atellica IM High-Sensitivity Troponin I (TnIH) result which was discordant relative to repeat testing when using TnIH lot 109. An initial elevated Atellica IM TnIH result was obtained for a 71-year-old male patient. A sample drawn one hour later produced much lower results, which were consistent across replicates. The initial sample was then re-tested twice, producing much lower results more consistent with those seen at one hour. An additional sample was drawn and tested, and it produced results consistent with those from the initial 1-hour sample. When these samples were later repeat-tested in triplicate on an Atellica CI analyzer, the results were all consistent with the lower observations from the Atellica IM instrument. The initial discordant, elevated result was never reproduced. There are no allegations of any adverse patient consequences in association with the observed result discordance. Atellica IM TnIH results, 2025-05-20: 0-hour sample: Sample ID 70607710 438.60 pg/mL 1:05PM – Elevated, discordant 33.56 pg/mL 2:40PM (rerun) 33.96 pg/mL 2:40PM (rerun) Rerun on CI at 3:16PM: 32.620 pg/mL 32.296 pg/mL 32.136 pg/mL 1-hour sample: Sample ID 70612710 27.23 2:18PM 28.22 2:38PM (rerun) 27.54 2:38PM (rerun) Rerun on CI at 3:16PM: Sample ID 70619210 26.428 pg/mL 26.299 pg/mL 26.677 pg/mL New Sample, tested 3:18PM: 27.81 pg/mL 27.79 pg/mL 27.07 pg/mL | yes | A customer from outside the United States reported observation of an elevated Atellica IM High-Sensitivity Troponin I (TnIH) result which was discordant relative to repeat testing when using TnIH lot 109. No issues were identified with assay calibration or Quality Control (QC). Siemens is investigating.
5 | 62315109 | 10997841 | 2026-01-15 | 2025-05-15 | other |  |  |  | No additional patient information was provided. | 2025-00121334-2：High Test Results | 2025-05-15 | medical organization | Kliniken Landkreis Heidenheim, GERMANY | The customer reports observation of an elevated Atellica IM High-Sensitivity Troponin I (TnIH) result which was discordant relative to repeat testing when using TnIH lot 109. An initial elevated Atellica IM TnIH result was obtained the patient in question. A sample drawn one hour later produced a much lower result, as did a two-hour sample. Re-testing of the initial sample produced a lower result consistent with the 1- and 2-hour sampling. There are no allegations of any adverse patient consequences in association with the observed result discordance. Atellica IM TnIH results, 2025-05-15: 0-hour sample Sample ID 70332210: 257.8 pg/mL – Elevated, discordant 1-hour sample Sample ID 70336310: 13.3 pg/mL 2-hour sample Sample ID 70343610: 15.4 pg/mL Re-test of initial sample: 12.09 pg/mL | yes | A customer from outside the United States reported observation of an elevated Atellica IM High-Sensitivity Troponin I (TnIH) result which was discordant relative to repeat testing when using TnIH lot 109. No issues were identified with assay calibration or Quality Control (QC). Siemens is investigating.
6 | 62315109 | 10997841 | 2026-01-15 | 2025-05-21 | other |  |  |  | No additional patient information was provided. | 2025-00121334-3：High Test Results | 2025-05-21 | medical organization | Kliniken Landkreis Heidenheim, GERMANY | The customer reports observation of an elevated Atellica IM High-Sensitivity Troponin I (TnIH) result which was discordant relative to repeat testing when using TnIH lot 109. An initial elevated Atellica IM TnIH result was obtained the patient in question. A sample drawn one hour later produced a much lower result, as did a two-hour sample. Re-testing of the initial sample produced a lower result consistent with the 1- and 2-hour sampling. There are no allegations of any adverse patient consequences in association with the observed result discordance. Atellica IM TnIH results, 2025-05-21: 0-hour sample Sample ID 70607710: 438.6 pg/mL – Elevated, discordant 1-hour sample Sample ID 70612710: 27.2 pg/mL 2-hour sample Sample ID 70619210: 27.8 pg/mL Re-test of initial sample: 32.6 pg/mL | yes | A customer from outside the United States reported observation of an elevated Atellica IM High-Sensitivity Troponin I (TnIH) result which was discordant relative to repeat testing when using TnIH lot 109. No issues were identified with assay calibration or Quality Control (QC). Siemens is investigating.