## Sheet1
| 医疗器械情况 | [9 empty] | 不良事件情况 | [11 empty] | 使用情况 | [5 empty] | 事件调查 |  | 评价结果 |  |  |  | 控制措施
ID* | 产品名称* | 注册证编号* | 产地* | 管理类别* | 产品类别* | 产品批号* | 产品编号* | UDI | 生产日期 (yyyy-MM-dd) | 有效期至 (yyyy-MM-dd) | 事件发生日期* (yyyy-MM-dd) | 发现或获知日期* (yyyy-MM-dd) | 伤害* | 伤害表现* | 姓名 | 出生日期 (yyyy-MM-dd) | 年龄单位 | 年龄 | 性别 | 病历号 | 既往病史 | 器械故障表现* | 预期治疗疾病或作用 | 器械使用日期* (yyyy-MM-dd) | 使用场所* | 场所名称* | 使用过程* | 合并用药/械情况说明 | 是否展开了调查* | 调查情况* | 关联性评价* | 事件原因分析* | 是否需要开展产品风险评价* | 计划提交时间 (yyyy-MM-dd) | 是否采取了控制措施* | 具体控制措施描述* | 未采取控制措施原因*
0 | [5 empty] | 481 | 10381206 |  |  | 2025-02-28 | 2025-03-02 |  | other | [5 empty] | Female |  | No other patient information provided. | 2025-00116604-3：High Test Results |  | 2025-03-02 | medical organization | Ireland Blyth Ltd Mauritius, MAURITIUS | The customer reported the observation of a falsely elevated IMMULITE 2000 HCG result using lot 481. The erroneous result was reported to the physician(s) and was questioned based on the clinical picture of the patient. The sample was repeated on a non-Siemens instrument and the lower result obtained was reported as the correct result to the physician(s). There are no known reports of patient intervention or adverse health consequences due to the discordant HCG results. Units: mIU/mL SID: Unknown (age unknown, female, not pregnant) Date of testing: 02-Mar-2025 Initial Result (erroneous): 33 Repeat Result obtained using a non-Siemens system (correct): 1.72 |  | yes | An Outside the United States (OUS) customer contacted a Siemens Remote Support Center to report an erroneous HCG patient result obtained when using lot 481 on an IMMULITE 2000 XPi instrument. The Limitations section of the IMMULITE 2000 HCG Instructions For Use (IFU) states: "For diagnostic purposes, the results obtained from this assay should always be used in combination with the clinical examination, patient medical history, and other findings." Siemens is investigating.
Note on G2: Author: 产品批号和编号不能同时为空
Note on H2: Author: 产品批号和编号不能同时为空
Note on J2: Author: 生产日期不能大于有效期至
Note on L2: Author: 事件发生日期不能大于当前日期,事件发生日期不能大于发现获知日期
Note on M2: Author: 发现或获知日期不能大于当前日期
Note on N2: Author: 伤害类型为严重伤害时，伤害表现不能为空！
Note on O2: Author: 伤害表现或器械故障表现不能同时为空！
Note on Q2: Author: 出生日期不能大于事件发生日期
Note on W2: Author: 伤害表现或器械故障表现不能同时为空！
Note on Y2: Author: 器械使用日期不能大于当前日期且不能大于事件发生日期
Note on Z2: Author: 使用场所为其他时,场所名称不能为空
Note on AD2: Author: 是否展开调查为是时,调查情况不能为空！
Note on AH2: Author: 展开产品风险评价时,计划提交时间不能为空
Note on AI2: Author: 计划提交时间大于当前日期
Note on AJ2: Author: 选择是需要填写具体控制措施描述,选择否需要填写未采取控制措施原因