## Sheet1
1 | 481 | 10381206 | 2025-02-28 | 2025-03-02 | other | Female | No other patient information provided. | 2025-00116604-3：High Test Results | 2025-03-02 | medical organization | Ireland Blyth Ltd Mauritius, MAURITIUS | The customer reported the observation of a falsely elevated IMMULITE 2000 HCG result using lot 481. The erroneous result was reported to the physician(s) and was questioned based on the clinical picture of the patient. The sample was repeated on a non-Siemens instrument and the lower result obtained was reported as the correct result to the physician(s). There are no known reports of patient intervention or adverse health consequences due to the discordant HCG results. Units: mIU/mL SID: Unknown (age unknown, female, not pregnant) Date of testing: 02-Mar-2025 Initial Result (erroneous): 33 Repeat Result obtained using a non-Siemens system (correct): 1.72 | yes | An Outside the United States (OUS) customer contacted a Siemens Remote Support Center to report an erroneous HCG patient result obtained when using lot 481 on an IMMULITE 2000 XPi instrument. The Limitations section of the IMMULITE 2000 HCG Instructions For Use (IFU) states: "For diagnostic purposes, the results obtained from this assay should always be used in combination with the clinical examination, patient medical history, and other findings." Siemens is investigating.